Wristwatch Personal Dosimeter - Specifications

1. Product Overview

The GM-16 Wristwatch Personal Dosimeter is an X and γ ray radiation warning instrument, primarily designed to monitor the radiation dose received by personnel from X-rays and γ-rays in radioactive environments. Featuring a compact size, it can be worn just like a standard wristwatch.

This dosimeter is suitable for measuring radiation dose equivalents in medical radiology, industrial non-destructive testing (NDT), radioisotope applications, occupational health prevention, military operations, nuclear power plants, and areas surrounding radioactive workplaces.

2. Key Features

High-Sensitivity Detector: High sensitivity with strong anti-interference capabilities.

Large LCD Display: Features extra-large fonts for easy reading and supports internal data storage.

User-Friendly: Easy to wear, simple to operate, and equipped with a combined data transfer and charging interface.

3. Technical Specifications

| Item / Parameter | Specification |
| --- | --- |
| Detector | Energy-compensated GM tube detector |
| Dose Equivalent Rate Range | Hp(0.07) 0.10 μSv/h to 30 mSv/h (Customizable upon request) |
| Cumulative Dose Equivalent | Hp(0.07) 0.01 μSv to 9999 mSv |
| Relative Intrinsic Error | ≤ ±15% (137Cs) |
| Energy Response Range | 50 keV to 1.3 MeV |
| Alarm Functions | Alarms for dose rate, cumulative dose, working time, etc. Continuously adjustable thresholds. |
| Data Transmission | Transfers measurement data to a computer via data cable |
| Power Consumption | < 2 mW |
| Operating Temperature & Humidity | Temperature: -10°C to +50°C Humidity: ≤ 95% (at 30°C, non-condensing) |
| Dimensions | Diameter 45 mm, Thickness 20 mm (excluding strap) |
| Weight | < 70 g (excluding strap) |